Stardust Calibration Log Sheet

Date	Time	Camera Temp	CCD Temp	Rad. Start	Rad. Finish	Rad. Filter	Sheet	
04/10/98		-8	-50.7				1	
Filter	Image Name	Exposure 
(sec)	Min DN	Mean DN	Max DN	Cal.
Lamp	Data
Comp.	
Nav Filter	145636	.02		2456		On	Off	
Nav Filter	145738	.02		2648		On	Off	
Nav Filter	145758	.02		2440		On	Off	
Nav Filter	145832	.02		2650		On	Off	
Nav Filter	145858	.02		2422		On	Off	
Nav Filter	145920	.02		2650		On	Off	
Nav Filter	150011	.00		502		On	Off	
Nav Filter	150032	.00		504		On	Off	

								

								

								

								

Date	Time	Camera Temp	CCD Temp	Rad. Start	Rad. Finish	Rad. Filter	Sheet	
04/10/98		-8	-50.7				2	
Filter	Image Name	Exposure 
(sec)	Min DN	Mean DN	Max DN	Cal.
Lamp	Data
Comp.	
C2	151336	20		502		Off	Off	
C2	151458	20		503		Off	Off	
C2	151548	20		506		Off	Off	
C2	150931	20		2079		On	Off	
C2	151141	20		2076		On	Off	
C2	151246	20		2082		On	Off	
C2	152008	20		176		On	On	
C2	152133	20		176		On	On	
C2	152216	20		176		On	On	
C2	151631	20		80		Off	On	
C2	151722	20		80		Off	On	
C2	151848	20		80		Off	On	
								
Yel Cont	152657	20		2612		On	Off	
Yel Cont	152757	20		Bad		On	Off	
Yel Cont	152902	20		2605		On	Off	
Yel Cont	152946	20		2622		On	Off	
Yel Cont	152450	20		198		On	On	
Yel Cont	152530	20		199		On	On	
Yel Cont	152610	20		198		On	On	
								
Red Cont	153837	1.5		186		On	On	
Red Cont	153858	1.5		187		On	On	
Red Cont	153920	1.5		186		On	On	
RedCont	154011	1.5		80		Off	On	

Date	Time	Camera Temp	CCD Temp	Rad. Start	Rad. Finish	Rad. Filter	Sheet	
04/10/98		-8	-50.7				3	
Filter	Image Name	Exposure 
(sec)	Min DN	Mean DN	Max DN	Cal.
Lamp	Data
Comp.	
RedCont	154032	1.5		80		Off	On	
RedCont	154052	1.5		80		Off	On	
RedCont	154408	2.0		2930		On	Off	
RedCont	154433	2.0		2935		On	Off	
RedCont	154458	2.0		2926		On	Off	
RedCont	154154	2.0		504		Off	Off	
RedCont	154337	2.0		502		Off	Off	
RedCont	154312	2.0		502		Off	Off	
								
IR Cont	154655	.1		2240		On	Off	
IR Cont	154732	.1		2277		On	Off	
IR Cont	154755	.1		2236		On	Off	
IR Cont	154824	.1		2272		On	Off	

IR Cont	154906	.1		500		Off	Off	
IR Cont	154929	.1		502		Off	Off	
IR Cont	154952	.1		504		Off	Off	
IR Cont	155031	.1		80		Off	On	
IR Cont	155105	.1		80		Off	On	
IR Cont	155133	.1		80		Off	On	
IR Cont	155156	.1		178		On	On	
IR Cont	155214	.1		180		On	On	
IR Cont	155235	.1		178		On	On	
IR Cont	155257	.1		179		On	On	

Date	Time	Camera Temp	CCD Temp	Rad. Start	Rad. Finish	Rad. Filter	Sheet	
04/10/98		-8	-50				4	
Filter	Image Name	Exposure 
(sec)	Min DN	Mean DN	Max DN	Cal.
Lamp	Data
Comp.	

								
NH2	155817	1.50		3006		On	Off	
NH2	155946	1.50		3007		On	Off	
NH2	160039	1.50		2986		On	Off	
NH2	160138	1.50		3006		On	Off	
NH2	160216	1.50		501		Off	Off	
NH2	160240	1.50		500		Off	Off	
NH2	160303	1.50		503		Off	Off	
NH2	155506	1.50		80		Off	On	
NH2	155528	1.50		80		Off	On	
NH2	155548	1.50		80		Off	On	
NH2	155633	1.50		216		On	On	
NH2	155653	1.50		214		On	On	
NH2	155729	1.50		215		On	On	
NH2	155749	1.50		214		On	On	
								
O1D	160522	2.30		2860		On	Off	
O1D	160555	2.30		2861		On	Off	
O1D	160624	2.30		2866		On	Off	
O1D	160652	2.30		2846		On	Off	
O1D	160848	2.30		503		Off	Off	
O1D	160733	2.30		500		Off	Off	
O1D	160758	2.30		502		Off	Off	

Date	Time	Camera Temp	CCD Temp	Rad. Start	Rad. Finish	Rad. Filter	Sheet	
04/10/98		-8	-50				5	
Filter	Image Name	Exposure 
(sec)	Min DN	Mean DN	Max DN	Cal.
Lamp	Data
Comp.	
O1D	160823	2.30		502		Off	Off	
O1D	160932	2.30		80		Off	On	
O1D	161025	2.30		80		Off	On	
O1D	161055	2.30		80		Off	On	
O1D	161116	2.30		80		Off	On	
O1D	161149	2.30		210		On	On	
O1D	161211	2.30		208		On	On	
O1D	161233	2.30		208		On	On	
O1D	161253			210		On	On	
								
Hi Res	161909	.25		3248		On	Off	
Hi Res	162010	.25		Bad		On	Off	
Hi Res	162053	.25		3248		On	Off	
Hi Res	162116	.25		3288		On	Off	
Hi Res	162140	.25		3249		On	Off	
Hi Res		.25				On	Off	
Hi Res	162234	.25		500		Off	Off	
Hi Res	162256	.25		502		Off	Off	
Hi Res	62319	.25		500		Off	Off	
Hi Res	161524	.25		80		Off	On	
Hi Res	161542	.25		80		Off	On	
Hi Res	161616	.25		80		Off	On	
Hi Res	161640	.25		80		Off	On	
Hi Res	161708	.25		225		On	On	
Hi Res	161727	.25		226		On	On	

Date	Time	Camera Temp	CCD Temp	Rad. Start	Rad. Finish	Rad. Filter	Sheet	
04/10/98		-8	-50.7				6	
Filter	Image Name	Exposure 
(sec)	Min DN	Mean DN	Max DN	Cal.
Lamp	Data
Comp.	
Hi Res	161746	.25		225		On	On	
Hi Res	161807	.25		226		On	On	
Hi Res	161827	.25		225		On	On	
Hi Res	161848	.25		226		On	On	
								
Yel Cont		20		504		Off	Off	
Yel Cont	153130	20		504		Off	Off	
Yel Cont	153209	.00		503		Off	Off	
Yel Cont	153237	.00		502		Off	Off	
Yel Cont	153352	.00		80		Off	On	
Yel Cont	153433	.00		80		Off	On	
Yel Cont	153526	20		80		Off	On	
Yel Cont	153610	20		80		Off	On